## Slide 1: Project Apollo
Why
  Ventilators need oxygen! (typical FiO2 0.3 … 1)
  Oxygen generation is a big problem in developing countries.
    No established infrastructure.
    Oxygen bottles are expensive
  People are already looking at alternative (local) ways for producing oxygen
What it is
  Goal of the Apollo prototype = enabling people around the world to build the prototype as fast as possible
  Focus = Simplicity and speed of build
    Open source, off-the-shelf materials
    Very low cost (aspirational target = $100 for 5 liters/min @ 90%)
  Final goal = Enable people to iterate and publish their own designs in the community
How to build it
  Follow the published build documentation
  Buy/source the materials (check out the BOM)
  Build the prototype
  Validate O2 concentration and flow. Use a good reference O2 and flow sensor for calibration
  Think about risk analysis and assessment: template for Apollo-derived design
  Document and iterate your own design. Publish your findings to the community!
Collaborations
  Helpful Engineering, Oxikit, Public Invention, Quick2space.org, Microsoft Garage
Documentation
  http://project-apollo.org

## Slide 2: Project Apollo – latest progress
Next version
  Focused on safety, user experience and maintainability
  O2 compatible materials in the oxygen path
  Dedicated PCB controller board, sensor integration, valve operational sensing
    3” TFT touch screen for diagnostic messages, medical-grade buzzer
    Auto-tuning (valve timing, auto-adjusts to changes in compressed air input pressure)
Status
  Mechanical design
    Focus on FDA approval and manufacturing
    Please see project “Oxygen Concentrator” from Public Invention (@Ben Coombs)
  Software prototypes for Apollo v4 – working
    Simulation software
      “device in the loop’ optimization strategies – tested and delivering close to theoretical
    Can control Apollo for O2 generation and basic optimization
    WiFi and Bluetooth integration TBD
  Sensor selection
    Sensors for pressure, humidity, temperature – selected and tested
    Valve board – schematic done, PCB in progress
    Controller board – schematic in review, PCB in progress
    Sensor board – schematic in progress, PCB in progress (similar with Apollo v3)

## Slide 3: Prototyping effort

## Slide 6: Conclusion
Apollo
  Open source oxygen concentrator with focus on low-cost
  When Apollo v4 comes out we encourage everyone to build one!
Please visit us at
  http://project-Apollo.org
    Github for designs, PCBs and software
  http://HelpfulEngineering.org
    Slack channel = #project-oxygen-concentrator